## Slide 1: District Heating Network
Audrey Blizard, Stephanie Stockar
11/24/2022
Manual for Data Acquisition Code

## Slide 2
Installation & Startup
  Installation
  Quick Start
  Input Files
  Main VI
Main VI
  LabVIEW Overview
  Structure
  Global Variables
  Adding additional components
Sub-Vis
  Intro
  Function
  Structure
Peltier Control
  Intro
  GUI Components
  Structure
Table of Contents

## Slide 3
  This code requires three programs to be installed on the operating computer
      NI LabVIEW 2020 is needed to run the data acquisition code and related sub-processes
        The LabView code collects data and processes it, providing both real-time control and visualization and the ability to save the collected data.
        It also allows the user to control the Peltier Junctions
        The ULX for NI LabView package must be installed for the code to operate correctly
      Microsoft Excel is needed to edit the Input, Calibration and Channels files
        .csv files are used to pass information into the code
        These files have a predesignated structure that must be followed for the code to operate correctly (see Section XX)
      InstaCal is needed to configure the Omega DAQs and set the calibration parameters for the thermistors
        The installation instructions for this software can be found in the ULX Quick Start Manual
Installation - Required Programs

## Slide 4
To interface with the Omega DAQ used to collect thermistor data, you must add the ULX for NI LabVIEW toolbox to the current LabVIEW version
Note that ULX for NI LabView is currently only compatible with LabVIEW 2020 and earlier.
This software can be downloaded from the Omega Website: https://www.omega.com/en-us/data-acquisition/data-acquisition-modules/omb-daq-2408-series/p/OMB-DAQ-2408
The Quick Start manual for this software can be found in the OneDrive folder “DistrictHeatingNetwork\Sensor Information\Thermistors”
ULX for NI LabVIEW also requires the installation of InstaCal to operate correctly
  Installation instructions can also be found the Quick Start Manual
  InstaCal is where you enter the Steinhart-Hart Constants used to calibrate the thermistors in the main LabVIEW VI
Installation - ULX for NI LabView

## Slide 5
Download the entire folder in lab_dhn
Required files
  Main Vi “DataAcquisitionCode”
  Global Variables file “globalvariables”
  14 sub-Vis used by the main VI. Details on these sub-vis can be found in Section 4
  5 files associated with the Peltier control, starting with “TC-48-20”
  3 .csv files used to pass information into the code
Installation - Program Files

## Slide 6
Open the main Vi “DataAcquisitionCode”
Ensure that only main file is open before starting the program
Use the White arrow located on the menu bar to start the program
At any time, the red circle can be used to abort operation of the program. However, the collected data will not be saved, and the Peltier controllers will continue running with the previously selected settings
Quick Start – Running the Program
Run
Abort

## Slide 7
Starting the program will open a pop-up window that will allow the user to configure the data collection settings
  Control Mode, Valve Control Mode, & Peltier Control Mode: Allows users to select how the code will operate, see next slide for details
  Time Between Samples (s): Sets the DAQ sampling rate
  Timer Length (s): Length of time for collecting data
  Number of Thermal Masses: Between 1-4
  Number of Sensors: Sensors being actively used, must be less than or equal to the number of sensors listed in the configuration file
  Number of USB DAQ: Number of connected Omega DAQs, defaults to 3
  Peltier 1 & Peltier 2: Communication ports for the Peltier Junction
  Folder Name: New folder to save data files, only required for control modes that save data
  Save Directory: Location for creation of new folder
Quick Start – User Inputs

## Slide 8
Control Mode
  Debugger (0): Allows user to visualize data without saving it. Only stops when aborted
  Manual (1): Allows user to visualize and save data in designated folder. Will only stop data acquisition manually, using the stop button
  Timed (2): Allows user to visualize and save data in designated folder. Stops data acquisition after a certain time has elapsed.
Actuator Control Modes
  Manual (Odd): Setpoints are entered manually by the user in real time using the main GUI
  Timed (Even): Setpoints are dictated by the file “Inputs.csv” and are entered as a function of iteration number
  Valve Control Mode
    Position (1, 2): Control valve position directly. Valve positions are entered manually by the user
    PID (3, 4): Uses PID controller to maintain a temperature setpoint in the thermal mass.
  Peltier Control Mode
    Temp (1, 2): Control the cold side temperature directly
    Power (3, 4): Uses desired power to calculate cold side temperature
    Tamb (5, 6): Uses desired simulated ambient temperature to set power and calculate cold side temperature
Quick Start – Operating Modes

## Slide 9
If a save mode is selected, the data is saved in the user defined location
Inside the “Save Directory” set by the user in the pop-up menu, a new folder with the name set by “Folder Name” will be created
Every time the code is run 4 .csv files will be created and written to inside this folder
  Peltier 1.csv: created by the Peltier 1 sub-vi, stores the Peltier 1 Sampling Time, Temp 1, Temp 2, Set Temp, and Power
  Peltier 2.csv: created by the Peltier 2 sub-vi, stores the Peltier 2 Sampling Time, Temp 1, Temp 2, Set Temp, and Power
  Processed.csv: Stores the processed data, converted from voltage to calibrated values, in the order specified by the “Channels.csv” file
  Raw.csv: Stores the original voltage values, in the same order specified by the “Channels.csv” file
Quick Start – Save Files

## Slide 10
Three excel files provide LabVIEW with configuration settings
  These files are loaded in programmatically and should be stored in the same folder as the main data acquisition code.
  The order of the rows in these files is used during code operation, so the structure of files must remain the same
  “Channels.csv”: Contains the information needed to collect data from each of the sensors and output it in the main Vi.
  “Calibration.csv”: Contains calibration coefficients for each of the sensor types
  “Inputs.csv”: Contains vectors of actuator control setpoint values with their corresponding execution iteration. Only used when either the valves or Peltier are set to Timed Control.
Input Files - Introduction

## Slide 11
Input Files - Channels.CSV
Provides the information need to create the channels in the sensor task
Also provides the headers used in the save files, ensure that the parameter names are valid MATLAB variable names
Additional information on the sensors can be found in the OneDrive folder "DistrictHeatingNetwork\Sensor Information\Sensor Placement Guide.docx”
Details on each line can be found on the next slide

## Slide 12
Parameter: Name given by the user to each sensor. This is only used in the output file and does not directly relate to GUI naming convention
Units: Units of calibrated sensor value, only used in output files
DAQ Channel: Channel in the DAQ. The structure of each channel name is as follows:
  USB DAQ: Dev#/Ti#, where the first number is the number assigned to the DAQ in InstaCal, and the second is the channel of the individual sensor
  TB-4353: PXI1Slot2/ai#, where the number is individual thermocouple
  BNC-2110: PXI1Slot3/ai#, where 0-15 are assigned to the first BNC block and 16-31 are assigned to the second
Sensor Type: TM-Thermistor, TC-Thermocouple, PS-Pressure sensor, MS-Mass Flow Sensor
Calibration Type: Indicates row to be used in the calibration.csv file
GUI Output #: Corresponds to the output fields in the main Vi. A table of these values can be found in the OneDrive folder "DistrictHeatingNetwork\Sensor Information\Sensor Locations.xlsx"
Upper V: Upper voltage limit the sensor can provide
Lower V: Lower voltage limit the sensor can provide
Input Files - Channels.CSV

## Slide 13
Provides information needed to convert collected voltage values into temperatures, mass flow rates, and pressure values.
Linear calibration with up to 9 terms:
Calibration values begin in row 4, with rows being references in “Channels.csv” file using the convention (i-3).
The calibration values are sorted by sensor type, with each starting at a different column
  TC: Column B
  PS: Column N
  MS: Column Z
  TM: N/A, as these sensors are calibrated via InstaCal
Details on the calibration values can be found in the OneDrive folder “DistrictHeatingNetwork\Sensor Information”
Input Files - Calibration.CSV

## Slide 14
Provides a vector of inputs to the actuators
Note that execution time is entered in iteration number, where
This vector must be sorted by iteration number, and can not have repeated iteration number values
Each actuator control mode is associated with its own set of rows
  Valve Timed-Position: Rows 3-4, enter the value position [%]
  Valve Timed-PID: Rows 7-8, enter the thermal mass temperature setpoint [C]
  Peltier Timed-Temp: Rows 5-6, enter the cold side temperature [C]
  Peltier Timed-Power: Rows 9-10, enter the desired peltier power [%]
  Peltier Timed-Tamb: Row 11, enter the simulated ambient temperature [C]
Values in the rows not associated with the selected actuator control mode are ignored
Input files - Inputs.CSV

## Slide 15
The main VI “DataAcquisistionCode” has three main sections on the GUI:
  Data Visualization pane (top right)
  Control pane (bottom right)
  Peltier pane (left)
The Data Visualization pane displays the collected data in real time
The control pane has two sections, one for the actuator controls, and one to control the data collection.
In the Peltier pane, each Peltier junction is controlled by a sub-VI, and the front panels are displayed in the main VI using subpanels
Main VI

## Slide 16
Has a built-in warning if the supply temperature exceeds 90°C
Each tabs provides data for 1 thermal mass, currently set up to accommodate up to 4 thermal masses
Thermal mass temperature is plotted as a sweep chart
Orange labels are TC and TM, purple are MS, and green are PS
Main VI- Data Visualization

## Slide 17
The data collection controls allow the user to see and control the operation of the code.
  # of Data Points Queued: Number of sets of data added to the queue to be stored after data collection ends
  Acquiring data: First phase of the code operation, storing data in the queue to be saved
  Stop Data Acquisition: Ends the first phase of the code operation, stops collecting data, and switches to saving it to the appropriate .csv files
  Processing Data: Second phase of the code, Storing data to the file
  # of Quick Restarts: Number of times user run an experiment with the same configuration. Appends a number to the save folder name
  Time Elapsed (s): Time elapsed in this run
  Time Limit (s): Available in timed code operation, shows the set timer length
Main VI- Control pane

## Slide 18
Based on which Valve control mode the user selects, a different control panels is visible
Main VI – Control Pane
Manual | PID
Either enter (Manual) or display (Timed) the desired valve positions using the dials. In Manual mode, it will not change valve position until the user clicks “Run” | In manual mode, user enters desired ThM setpoints in the top box. In timed mode, the setpoints are displayed here. The calculated valve position is displayed below. The user can also set the PID gains. The update rate dictates how often a new control command is sent to the valves.

## Slide 19
Based on which Peltier control mode the user selects, a different control panels is visible
Main VI – Control Pane
Cold Side | Power
Either enter (Manual) or displays (Timed) the desired Peltier cold side temperature using the sliders. In Manual mode, it not update the Peltier controller until the user clicks “Go” | In manual mode, user enters the desired power in the top box. In timed mode, the set powers are displayed here. In Tamb mode, the simulated ambient temperature is entered/displayed in the bottom box. The calculated necessary cold side temperature is displayed below. The update rate dictates how often a new control command is sent to the Peltier controller.

## Slide 20
Control Sensor: Thermocouple located on the cold side of the Peltier junction
Setpoint: Desired cold side temperature
Input 2 Sensor: Thermocouple located at the top of the ThM
Run: Turns the Peltier junction on. IMPORTANT: Stop running before ending data collection
Power: % Power output
Array #: Array location where the Peltier control variables are saved during the code operation
COM: Communication port used by the Peltier
Main VI – Peltier Pane

## Slide 21
NI LabVIEW is a block diagram-based coding language use to interface with the NI DAQ
The front panel acts as the display for the code. All outputted and inputted values are shown here in type-specific display variables
The block diagram is where the code is written
  LabVIEW uses type specific variables, and each type has a different wire style
  Arrays and clusters can both be used to hold groups of variables, where arrays hold variables of the same type, and clusters hold variables of different types
  The code attempts to run each section as soon as all required inputs are received. Therefore, it may be necessary to wire inputs into a sequence structure or a case structure to ensure it does not execute too early. Error variables can also be used enforce order
Sub-Vis are custom blocks that can be inserted into a block diagram to call a set sub function. A sub-vi can not be called while it is already running. They are used to help keep a code compact and easy to follow. Variables on the front panels of these vis can be set to be output and input by the associated block
Global variables are used to pass information between Vis. It can be difficult to tell execution order in these variables, so make sure they are written before they are read. Global variables are particularly useful for passing values into sub-vis that are already running, and need values updated, since inputs are only taken once, when the sub-vi is first called
LabVIEW is a zero indexing program, meaning the first element is enumerated as 0. This is true in both loop and array indexing
Some sub-vis are configured as subroutines. These allow for faster code execution, at the expense of being able to probe the values during troubleshooting. This can be change by editing File -> VI Properties -> Execution -> Priority
LABVIEW Overview

## Slide 22
Structure
While Loop 1 allows for repeating experiments once the code is started. It is exited once the user selects “No” when prompted to “Perform another test?”
For Loop 1 set to execute once. Loads the calibration parameters
While Loop 2 allows for quick restarts, where the code is restarted without changing any of the settings. It executes after For Loop 1 finishes. This loop creates the save files, queues, and tasks. Exits when the user selects “No” when asked “Would you like to Quick Restart?”
While Loop 3 starts the two Peltier control sub-vis and contains the data collection and data processing processes. Exits once the data processing is complete.
While Loop 4 controls data collection. Repeats at every time step to collect data, at the frequency set by the user. Exits when an error occurs, when the timer expires in the timed mode, or when the user presses the “stop data acquisition” button.
Sequence Structure 2 controls data processing. The first panel saves the data if there are no errors and the second shuts down the Peltier sub-vis. Exits once the data is saved and the Peltier control sub-vis are shut down.
Sequence Structure 1 gathers the user inputs and then configures the front panel based on the user selected inputs.

## Slide 23
Structure
While Loop 1
For Loop 1
Load Calibration parameters from file
Displays message to user and either restarts or terminates this loop

## Slide 24
Structure
Controls the number of visible display tabs in the main vi, based on the number of ThM
Initializes the global variable used to store the current Peltier control sensor values
Initializes the global variable used to tell the Peltier sub-vis a change has been made to the control values
Changes the control pane display based on the valve control mode
Changes the control pane display based on the Peltier control mode
Hides some of the control pane components if debug mode is chosen
sub_get user input.vi
sub_process channels.vi
Sequence Structure 1

## Slide 25
Structure
sub_create tasks.vi
Creates queues to hold the raw and processed data before it is saved
Creates save files if debug mode is not selected
Displays message to user and either restarts or terminates this loop
Deletes the queue after the data collection is over
sub_create save files.vi
Ends all the tasks used to collect data from the Omega DAQs
Ends the task used to collect data from the NI DAQ
Ends the task used to control the valves
While Loop 2
Loads the inputs file
Initializes the shift register that stores the ambient and thermal mass temperatures for use in the actuator control

## Slide 26
Structure
While Loop 3
sub_peltier 1 control.vi
Finds the appropriate subpanels in the Peltier Pane and either inserts the matching subvi or makes it invisible based on the number of thermal masses
sub_peltier 2 control.vi

## Slide 27
Structure
sub_run actuators.vi
sub_collect data.vi
Checks for error, by looking for a large amount of data collected immediately
Ends loop if 1) button is pressed, 2) time elapses, or 3&4) an error occurs
Displays time limit and elapsed time & calculate if time has elapsed
Displays current values in the GUI
sub_process data.vi
Shift register stores which command column to execute
While Loop 4 Part 1
Processes data, and when not in debug mode, adds data to the queue

## Slide 28
Structure
Converts update rates from seconds to iterations
Stores PID gains in global variable
Stores “Run” & “Go” Button actions in global variables
Stores and displays Peltier values based on which control mode is selected
Stores and displays Peltier values based on which control mode is selected
Stores ambient and thermal mass temperatures for use in the actuator control
Displays warning if supply temp exceeds 90°C
Changes control pane display settings
While Loop 4 Part 2

## Slide 29
Structure
Displays if an error occurred
Removes the Peltier sub-vis from the subpanels
sub_save output files.vi
Ends While Loop 3
Saves data if no error occurred
Stops saving Peltier values
Sequence Structure 2
Changes control pane display settings
Ends the Peltier controller sub-vis

## Slide 30
Global Variables
Name | Type | Description | Write location
# of sensors | Long | # of rows to be taken from Channels.csv | sub_acquire user input
# of ThM | Long | # of Thermal Masses | sub_acquire user input
# of USB DAQ | Long | # of Omega DAQs to be added to task | sub_acquire user input
Control Mode | Long | Control mode (see slide XX) | sub_acquire user input
Timer Length | Double | Timer length for timed mode | sub_acquire user input
File Name | String | Name for save folder | sub_acquire user input
Path | File Path | Path for new save folder | sub_acquire user input
Sensor Descriptions | 2D array of String | Parameter, Unit, DAQ Channel, & Sensor Type taken from Channels.csv | sub_process channels
Sensor Labels | 2D array of Long | Calibration Type and GUI Output # taken from Channels.csv | sub_process channels
Voltage Limits | 2D array of Double | Upper and Lower Voltage Limits taken from Channels.csv | sub_process channels
Sources | 1D array of Long | DAQ Type for each sensor, 1 for Omega, 0 for NI | sub_create tasks

## Slide 31
Global Variables
Name | Type | Description | Write location
Calibration Table | 2D array of Double | Calibration values taken from Calibration.csv | For Loop 1
Valve Control Mode | Word | Valve control mode (see slide XX) | sub_acquire user input
Valve New Value | Boolean | Update Valve, from “Run” button | Main GUI
Valve Position | 1D array of Double | Desired Valve position (%) | Main sub_run actuators
Valve Set Temp | 1D array of Double | ThM Temperature setpoints for PID Valve Control | Main GUI sub_run actuators
Valve Gains | 1 D array of Cluster of 3 Double | Gains for PID control of Valves | Main GUI
Valve Update Rate | Long | How often the valve position is updated | Main GUI

## Slide 32
Global Variables
Name | Type | Description | Write location
Peltier Control Mode | Word | Peltier control mode (see slide XX) | sub_acquire user input
Peltier New Value | Boolean | Update Peltiers, from “Go” button | Main GUI
Peltier Set Temp | 1D array of Double | Cold side temperature setpoints for Peltier Junctions | Main GUI sub_run actuators
Peltier Set Power | 1D array of Double | Power setpoints for Peltier Junctions | Main GUI sub_run actuators
Peltier Control Sensor | 1D array of Double | Current cold side temperature measurement | Main GUI sub_peltier control
Peltier Update Rate | long | How often the Peltier controller is updated | Main GUI

## Slide 33
Global Variables
Name | Type | Description | Write location
Peltier Save | Boolean | Save Peltier data to file | Main GUI sub_peltier control
Peltier Stop | Boolean | End execution of the Peltier sub-vis | Main GUI
Peltier Update | 1D array of Boolean | Indicates when a Peltier global variable value has changed so the Peltier control sub-vis can check them | Main GUI sub_run actuators
Peltier Tamb | Double | Simulated ambient temperature for Tamb control of Peltiers | Main GUI Sub_run actuators

## Slide 34
Currently, the code is configured to accommodate 4 thermal masses, 2 Peltier junctions, and 2 control valves
While it easy to add the sensors for the 3rd and 4th thermal mass (see below), follow the instructions to add the valves and Peltier junctions associated with these thermal masses
Adding a 3rd & 4th Thermal Mass
  GUI: The code is already setup to display data for a 3rd & 4th thermal mass, the only required change is to rearrange the tabs used to display the data. The ThM 3 & ThM 4 tab can be placed in between the Source tab and the ThM 2 tab.
  Block Diagram: No changes needed.
  Sensor: Add the sensor to the channels file. Use the file “DistrictHeatingNetwork/Sensor Information/Sensor Locations.xlsx” to fill in the GUI Output #.
Additional instructions are also included to add thermal masses beyond the 4 the code is currently configured to handle.
Adding Additional Components

## Slide 35
GUI:
  Duplicate the ThM 2 tab, ensure the new tab control has a name of the form “ThM #”, and all variables.
  Rearrange the tabs so they match the lab’s configuration. ThM 1 should always be the last display as that tab contains fewer sensors and a different pipe diagram.
Main Block Diagram:
  Displaying the data: Adding the new tab will create terminals in the block diagram. Add additional cases the to the case structure used display current values in the GUI in While Loop 4. Put the variables in new cases, recording the variable/case # pairs in the file “DistrictHeatingNetwork/Sensor Information/Sensor Locations.xlsx.”
Sub Vis:
  sub_get user input: increase the upper limit on the number of thermal masses error to the appropriate value
Sensors:
  Use the variable/case # pairs to populated the GUI Output # in the Channels file
  Each additional thermal mass requires 8 TM, 4 PS, and 2 MS to be fully observable
Adding Additional Thermal Masses

## Slide 36
GUI:
  Add additional subpanels to the left pane in the main GUI. Make sure it is labeled “Peltier Control #”.
  Peltier Control Menu tab control:
    Manual tab: Add additional elements to the control array
    Timed tab: Add additional elements to the indicator array
    Power tab: Add additional elements to the Power Setpoint control array. Add additional elements to the Control Temp indicator array.
Main Block Diagram: No changes needed
Sub Vis:
  Sub_peltier control #: Create 1 copy of “sub_peltier control 1.vi” for each new Peltier junction and name it “sub_peltier control #.vi”, with the number correlating to its respective thermal mass.
  Sub_run actuators.vi: Additional rows will have to be added to the Inputs.csv file for each additional Peltier junction. Doing this will change the indexing for the control inputs. There are 6 total to update. Each corresponds to a different timed control mode and are found in the odd # cases in both structures.
Sensors: Currently the Peltier controller is configured to have 2 10K control thermistors. This can be changed manually by changing the default values in the front panel of the new sub vi.
Inputs.csv: Add additional rows to the file, keeping them grouped by input type. Add one row for each Peltier Junction to the “Cold Side Temp”, and “Power” sections. Only 1 Tamb row is needed for all Peltier junctions. Additionally, add more rows for the heat transfer coefficients of each thermal mass.
Adding Peltier Junctions

## Slide 37
GUI:
  Valve Control Menu tab control:
    Manual tab: Add additional elements to the control array
    Timed tab: Add additional elements to the indicator array
    Power tab: Add additional elements to the Temp Setpoint control array, Valve Position indicator array, and PID cluster array.
Main Block Diagram
  If more than 4 valves are being used, add additional cases to the case structure in While Loop 4 use in the PID control of the valves. Add local variables referencing the appropriate thermal mass temperatures.
Sub Vis:
  Sub_create tasks.vi: add additional channels for each new valve, in the appropriate form, separated by commas to the I/O constant used to populate the physical channels of the actuator task.
  Sub_run actuators.vi: Additional rows will have to be added to the Inputs.csv file for each additional valve. Doing this will change the indexing for the control inputs. There are 6 total to update. Each corresponds to a different timed control mode and are found in the odd # cases in both structures.
Inputs.csv: Add additional rows to the file, keeping them grouped by input type. Add one row for each valve to the “Position”, and “PID Setpoint” sections.
Adding Control Valves

## Slide 38
There are 11 custom sub-Vis created to perform various operations during the code’s execution. Their functions and structure will be presented in this section.
The 11 sub-Vis are:
  sub_calibrate sensor data.vi
  sub_collect data.vi
  sub_ create save files.vi
  sub_create tasks.vi
  sub_get user input.vi
  sub_ omega tasks.vi
  sub_ process channels.vi
  sub_ process data.vi
  sub_ process data no queue.vi
  sub_ run actuators.vi
  sub_ save output files.vi
Summary

## Slide 39
sub_calibrate sensor data.vi | 
Icon: | 
Description: | 
Converts sensor data from voltage to calibrated values, based on the Calibration.csv file | 
Variables: | 
Inputs: Raw Data: Raw data from a single sensor Sensor Column: Column in the Channels.csv file for the sensor | Outputs: Calibrated Data: Data converted from voltage to appropriate value
Function
sub_collect data.vi | 
Icon: | 
Description: | 
Gets the current sensor readings from the DAQ systems and the current time | 
Variables: | 
Inputs: Sensor Task: Task for NI DAQ Sensor Error: Error for NI DAQ Omega task: array of tasks for Omega DAQs Omega error: array of errors for Omega DAQs Iteration: iteration # of While Loop 4 | Outputs: Sensor Data: array of raw voltage data from sensors Elapsed Time (s): Time when data was collected All Initialized?: True if expected number of channels exists

## Slide 40
sub_ create save files.vi | 
Icon: | 
Description: | 
Creates the files where the data will eventually be saved | 
Variables: | 
Inputs: # of Quick Restarts: number to be appended to file name | Outputs: Save Path: Folder where the data will be stored Peltier Save Files: File names for the created Peltier data files Error out: error from creating these files and folders
Function
sub_create tasks.vi | 
Icon: | 
Description: | 
Creates the tasks used to get the data from the DAQ systems | 
Variables: | 
Inputs: DAQ Sample Rate: Frequency for DAQ Sampling | Outputs: Output cluster: cluster of tasks and errors from the NI and Omega DAQ

## Slide 41
sub_get user input.vi | 
Icon: | 
Description: | 
Gathers users input needed to control various aspects of the data collection | 
Variables: | 
Inputs: None | Outputs: DAQ Sample Rate: Inverse of samples per second Peltier Channels: Array of COM port numbers for Peltier junctions
Function
sub_ omega tasks.vi: | 
Icon: | 
Description: | 
Creates the tasks used to get the data from the appropriate Omega DAQ | 
Variables: | 
Inputs: DAQ Channel: channel for the current sensor Omega tasks: array of Omega tasks Omega errors: array of errors associated with Omega Tasks | Outputs: Source Number: # of Omega task for current sensor (1 indexed) Omega tasks out: new array of Omega tasks Omega errors out: new array of errors associated with Omega Tasks

## Slide 42
sub_ process channels.vi | 
Icon: | 
Description: | 
Processes the Channels.csv file into global variables | 
Variables: | 
Inputs: None | Outputs: None
Function
sub_ process data.vi | 
Icon: | 
Description: | 
Takes the raw data, processes it, and adds it to the appropriate queues | 
Variables: | 
Inputs: Elapsed Time: time when data was collected Processed Queue in: queue for processed data Processed error in: error associated with the processed queue Raw Queue in: queue for raw data Raw error in: error associated with the raw queue Raw Data in: 1D array of raw sensor data | Outputs: Processed Queue out: new queue for processed data Processed error out: new error associated with the processed queue Raw Queue out: new queue for raw data Raw error out: new error associated with the raw queue Processed Data Out: 1D array of processed sensor data

## Slide 43
sub_ process data no queue.vi | 
Icon: | 
Description: | 
Takes the raw data and processes it without adding it to the queues | 
Variables: | 
Inputs: Elapsed Time: time when data was collected Raw Data: 1D array of raw sensor data | Outputs: Processed Data: 1D array of processed sensor data
Function
sub_ run actuators.vi | 
Icon: | 
Description: | 
Performs the desired actuator tasks based on the current iteration and control mode | 
Variables: | 
Inputs: Elapsed Time: time when data was collected Raw Data: 1D array of raw sensor data | Outputs: Processed Data: 1D array of processed sensor data

## Slide 44
Function
sub_ save output files.vi | 
Icon: | 
Description: | 
Takes the data from the queues and saves it to the appropriate .csv files | 
Variables: | 
Inputs: Raw Queue in: queue for raw data Processed Queue in: queue for processed data Processed Data error: error associated with the processed queue Raw Data error: error associated with the raw queue | Outputs: error out: combined error from saving data

## Slide 45
Structure
Global Variables with Calibrate.csv data
Input Variables
Check the sensor type, TM don’t need calibration
Starting column in Calibrate.csv
Sum up terms
Global Variables with Channels.csv data
Row in Calibrate.csv from Channels.csv
Loop through 10 terms
Each loop, raise voltage to increasing power, and increase coefficient column
Output
sub_calibrate sensor data

## Slide 46
Structure
Intermingles sensor data from the 2 types of DAQs, where 0 is NI, > 0 is Omega
Reads the NI DAQ sensors
Reads the Omega DAQs sensors
Gets the time when the data collection is completed
Ensures all data is collected before the time is recorded
sub_collect data

## Slide 47
Structure
Adds a number to the folder name if a quick restart is requested
Adds columns for the time and valve positions to the columns created in the channels file
Creates the correct number of Peltier save files
Writes the column headers to the two save files (raw and processed)
Creates the folder
sub_ create save files

## Slide 48
Structure
Creates tasks for the DAQs
Creates output channels for the two valves
Sets up the same reference clock for both NI tasks
Provides the channel settings
Adds the channel to the appropriate task based on sensor type
Vector of which DAQs each channel belongs to
Clusters all tasks and errors together
Sets no clock for the Omega tasks, since the NI DAQ handles all the timing
Starts the sensor tasks
Case Structure contains sub_omega tasks.vi to assign the channels for the type “TM”
sub_create tasks

## Slide 49
Structure
User inputs, Details can be found in “Quick Start – User Inputs”
While Loop, waits for user to click “Apply Options” before proceeding.
While Loop, exits only if none of the errors are triggered.
Error checks
Some inputs are not required in certain control modes
Stores variables in global variables & indicators
sub_get user input

## Slide 50
Structure
Extracts the DAQ # from the channel name
Gets the correct task and error
Creates the channel
Adds the new task and error back in
sub_ omega tasks

## Slide 51
Structure
Warning Messages to user, code keeps running
Path for the Channels file
Loads data with different types specified
Stores components in global variables
Error checks, purpose explained in corresponding warnings
sub_ process channels

## Slide 52
Structure
sub_calibrate sensor data.vi
Appends valve position and time to processed data
Appends time to raw data
Processes all elements of raw data array
Adds data to appropriate queue
sub_ process data

## Slide 53
Structure
Processes all elements of raw data array
Ensures time is available before beginning processing
sub_calibrate sensor data.vi
sub_ process data no queue

## Slide 54
Structure
Valve control case structure
Peltier control case structure
Ambient and Thermal Mass temperatures
Thermal mass characteristics
Passes calculated control input to Peltier controllers
Sends voltage command to valves task
Selects cases based on user inputted control modes
Converts from power to cold side temperature if true
sub_ run actuators

## Slide 55
Structure
Takes manually entered values directly
PID with setpoints entered manually
Checks if the valve control should update
Takes position from inputs file
Increments counter if a new input is ran
Updates valve position if iteration # matches input file
PID with setpoints from input file
Updates PID setpoint if iteration # match, otherwise uses old value
Sub_Run Actuators (Valve Cases)

## Slide 56
Structure
Takes manually entered cold side values directly
Takes manually entered power values directly
Checks if the Peltier control should update
Takes cold side temp from inputs file
Updates Peltier if iteration # matches input file
Calculates power from manually entered simulate ambient temperature
Updates power setpoints from inputs file if iteration # match, otherwise uses old values
Calculates power from simulate ambient temperature, either from inputs file, if iteration # matches, or from old value
Sub_Run Actuators (Peltier Cases)

## Slide 57
Structure
Flushes all data from the queues
Removes external cluster holder from 1D array of data, skipping the 1st value
Appends data to the pre-created files, given by the global variable path
sub_ save output files

## Slide 58
The Peltier controllers are communicated with using the sub vis “sub_peltier # control” where each sub vi interacts with one Peltier controller. These vis pass information to the controller and receives information from the controller and displays it. It also records the data.
To operate the controllers, the project files and library of subvis must be in the same folder as the main sub vi.
To ensure the GUI displays properly in the main vi, shrink the window to fit the Display tab, and save it centered around this tab
Introduction

## Slide 59
This code was modified from the code provided by TE Technology. The main changes are:
  The code was simplified, cleaned up, and redundant elements were removed. This makes troubleshooting easier and adapts the code to run as a subvi.
  The main GUI was changed to make it easier to display in the subpanel.
  The control cluster was changed to add the End and Final Loop variables, and some extraneous variables were removed.
  The Stop and Quit LabVEIW commands were also removed in the original code, replaced by a variable that will end the main for loop and allow the code to stop running.
  The Stop and Quit LabVIEW commands were also eliminated in the sub vis. They were replaced with outputs that could be used to modify the End variable’s state
  Global variables were added to allow real time communication with the main VI.
  All hidden variables on the main vi were either removed or made visible.
  The ability to download configurations was removed, as this was poorly implemented and not used on our experiments.
  Additional information was added to the error messages to make troubleshooting easier
Introduction

## Slide 60
The GUI has 5 tab controls that group the variables. The categories are:
  Display: Displays relevant data to the user.
  Misc Settings: Settings and outputs from the code. None of these should be modified unless, changes are made to the Peltier junction.
  Graph Data: Data stored to be used in the graph and to set the Peltier controller.
  Alarm Settings: Settings to configure alarms. No alarms have been used, and this code has not been tested. See manual for details on the variables meanings and operation.
  Controller Settings: Settings for how the power is calculated based on the set cold side temperature. These must be set differently, depending on the Peltier control mode See manual for more details on how the PID control is implemented.
GUI Components

## Slide 61
GUI Components- Misc Settings
Name | Type | Description | Default Value
Sensor Type | Radio Button | Type of thermistor connected to the Peltier controller | 10K
Output Polarity | Radio Button | Configures the Peltier junction to either heat or cool | Cool
Temperature 2 Display | Radio Button | Indicates if a second thermistors connected to the Peltier controller | On
CRC Error | Boolean | Indicates if a CRC error occurs during code execution | -
EEPROM Write Enable | Boolean | If true, the controller remembers its settings even after power to the controller has been turned off. See manual for more details | False
Firmware Version | String | Displays the firmware version read from the controller | -
Analog Output Multiplier | Double | Scales the maximum voltage that can be supplied to the Peltier Junction | 1
Reading Configuration | Single | Displays the progress of the configuration of the Peltier controller as a % | -
Error out | Error | Displays the error cluster as the code runs | -

## Slide 62
GUI Components- Graph Data
Name | Type | Description
Set Temperature | Double | Set temperature to be communicated to the controller
Control Sensor | Double | Displays Sensor 1 value
Input 2 Sensor | Double | Displays Sensor 2 value
Elapsed Time | Double | Start time minus current time, used when storing the data in the save file
Sample in Time | Time Stamp | Current time
Start Sample Time | Time Stamp | When the data started being recorded

## Slide 63
GUI Components- Control Settings
Name | Type | Description | Power Control Value | Cold Side Control Value
Set Limit High | Double | Maximum possible set temperature | 100 | 70
Set Limit Low | Double | Minimum possible set temperature | -20 | -20
Control Sensor Offset | Double | Offsets the control sensor reading by this amount for display and control calculations | 0 | 0
Proportional Bandwidth | Double | Temperature span over which the power is proportioned from 100% to 0% power, centered about the temperature set point. | 50 | 5
Integral Gain | Double | Shifts the proportional bandwidth to correct for offsets | 0 | 1
Derivative Gain | Double | Shifts the proportional bandwidth based on the change in the error | 0 | 0
Sample Rate (s) | Integer | Sets how often the data is collected | 1 | 1

## Slide 64
Structure
While Loop 1 sets up communication with controller based on COM variable. Changes end to true if an error occurs.
Error 1 displays error message and restarts While Loop 1 if user selects “restart”, ends it if user selects “exit,” or no error occurs
While Loop 2 executes while the Peltier controller is running. It updates Peltier variables if values of the relevant global variables change. Exits if the End variable is true. Updates the values of the control cluster base on outputs from Case Structure 2. Displays the error variable and the CRC error Boolean.
Sequence Structure 1 initializes connection with Peltier controller.
Error 2 displays an error message if an error occurs during Sequence Structure 1. Either sets the control cluster to restart or end.
Sequence Structure 2 performs ordered tasks associated with the continued operation of the Peltier controller. Contained in an error case, so only executes if no error is present
Pane 1 tries to communicate with the Peltier controller. Executes if Read Model is true. Gets and displays the firmware version. Ends loop if the controller connects, if a CRC error is detected, or if it tries more than 3 times. Contained in an error case, so only executes if no error is present
Pane 2 loops through settings to initialize them. Executes if Read Settings is true. Ends once all 21 settings have been configured. Contained in an error case, so only executes if no error is present
Pane 1 checks alarm status and blinks Alarm light if the alarm value is in range. Set Read Settings to true if necessary.
Pane 2 reads the cold side temperature, inputs 1 & 2 (if necessary), displays them and puts them in an array.
Pane 3 displays the values in the strip chart if no error has occurred. Reads the current time.
Pane 4 reads and displays the power output and adds it to the data array.
Pane 5 saves the data to the file. Creates the header if the file is empty. Calculates elapsed time.
Case Structure 2 reads the set temperature from the controller and Input 2 if set to.
Case Structure 3 reads Input 1 sensor and saves it as in a global variable.
Error 4 displays an error message if an error occurs during While Loop 2. Either sets the control cluster to restart or end.
Event Structure 1 updates the controller if any changes are made to the setting. Contained in an error case, so only executes if no error is present
Case Structure 1 displays error message if CRC Error is true and either sets the control cluster to restart or end. Otherwise passes the control cluster unchanged.
Error 3 displays an error message if an error occurs during While Loop 2. Either sets the control cluster to restart or end.

## Slide 65
Structure
While Loop 1
Control cluster that holds variables to indicate what tasks must be performed
Sets communication port #
Initializes error
Connects with Peltier controller port
Error 1 pauses then re-runs loop if an error occurred and user selected “restart”. Sets end to true if user selects “Exit.” Ends loop if no error occurred
Sets end to true if COM port number is zero or negative

## Slide 66
Structure
While Loop 2
Resets the CRC Error value
If a global variable is updated, the updated value is assigned to the appropriate local variable
If the global variable says to end the code, the run Peltier button is set to false to turn of the controller power
Once the update is applied, need to update variable is set to false
Values stored in arrays so array # is used to index

## Slide 67
Structure
Sequence Structure 1 Pane 1
Executes if Read Model is true
Does not execute if an error occurs
Reads information from the controller
Clears strip chart history
Gets and displays firmware version from controller
Tries to connect 3 times before ending the codes execution
Ends loop if it connects correctly, a CRC error occurs, or it tries 3 times

## Slide 68
Structure
Sequence Structure 1 Pane 2
Executes if Read Settings is true
Does not execute if an error occurs
Loops through 21 settings
Displays progress of configuration
Changes if a CRC error occurs or an error occurs that requires ending the execution
Settings have now been read
Performs action and saves value in appropriate local variable

## Slide 69
Structure
Error 2
If End is true, code needs to loop through 1 more time to ensure the controller is turned off, if possible. The Final Loop variable ensures this happens
Sets code to reconfigure model and settings if an error occurred and user selected “restart”. Sets end to true if user selects “Exit.” Does nothing if no error occurred.

## Slide 70
Structure
Sequence Structure 2 Pane 1
Displays if an alarm has occurred
Recheck settings
Changes if a CRC error occurs or an error occurs that requires ending the execution
Checks the controller alarms

## Slide 71
Structure
Sequence Structure 2 Pane 2
Changes if a CRC error occurred or an error occurred that requires ending the execution
Changes if a CRC error occurred
Stores control sensor value in global variable
Reads Sensor 1 value
Stores all values in an array
Turns on sampling indicator
Reads Sensor 2 value, if available
Reads and displays setpoint value
Only runs if no error has occurred

## Slide 72
Structure
Sequence Structure 2 Pane 3 & 4
Gets current time
Displays sensor values on strip chart, based on how many sensors are available
Sets chart limits and labels
Only runs if no error has occurred
Reads output power from controller
Displays output power
Turns off sampling indicator
Adds power to array of values

## Slide 73
Structure
Sequence Structure 2 Pane 5
Processes current time to elapsed time. This is reset if the code is not saving data and on the first loop
Only executes if data should be saved
Sets code to reconfigure model and settings if an error occurred and user selected “restart”. Sets end to true if user selects “Exit.” Does nothing if no error occurred.
Saves array to file
Adds headers if file is empty
Only executes if no error has occurred reading the file size
Path is an input of the sub vi

## Slide 74
Structure
Event Structure 1
Checks for changes in values of control variables. If a change occurs, the event structure communicates the new value to the controller and then performs checks to ensure the value is implemented correctly an no errors/alarms have occurred.
Outputs any CRC errors or End loop commands
No CRC error if the code cannot execute due to an error
Most cases have error checks, where the code will not execute if an error has occurred
Sample rate (ms) is used to hold the event structure and control rate of looping. The event structure is held until it is time to increment the data, then a timeout occurs, and the loop continues

## Slide 75
Structure
Error 3 & Case Structure 1
Changes the CRC variable if an error occurred in the event structure
Changes the End variable if an error occurred in the event structure
Sets code to reconfigure model and settings if an error occurred and user selected “restart”. Sets end to true if user selects “Exit.” Does nothing if no error occurred.
Exits loop if end is ever set to true and the code has had the opportunity to loop through the event structure and stop the controller
Sets code to reconfigure model and settings if a CRC error occurred and user selected “restart”. Sets end to true if user selects “Exit.” Does nothing if no error occurred.